## Slide 1: BMS 524 - “Introduction to Confocal Microscopy and Image Analysis”
The Principles of Microscopy: II
Department of Basic Medical Sciences,
School of Veterinary Medicine
Weldon School of Biomedical Engineering
Purdue University
J. Paul Robinson, Ph.D.
SVM Professor of Cytomics & Professor of Biomedical Engineering
Director, Purdue University Cytometry Laboratories, Purdue University
This lecture was last updated in January, 2021
This PowerPoint lecture is available at http://tinyurl.com/2dr5p
Find other PUCL Educational Materials at http://www.cyto.purdue.edu/class
These slides are intended for use in a lecture series. Copies of the slides are distributed and students encouraged to take their notes on these graphics. All material copyright J. Paul Robinson unless otherwise stated. No reproduction of this material is permitted without the written permission of J. Paul Robinson. Except that our materials may be used in not-for-profit educational institutions with appropriate acknowledgement. It is illegal to post these lectures on Course Hero or any other site.
Note: AUDIO is available
In this lecture by clicking
On the speaker icon on
each slide

## Slide 2: Introduction to Lecture 2
Principles of Microscopy II
Magnification
Nature of Light
Basic Microscopes -Optical Designs
Numerical Aperture
Refractive Index

## Slide 3: Learning Objectives – Lecture 2
Understand the different type of microscopes and what their different applications are
Learn the basics of the most important things that define restrictions on optical detection systems
The electromagnetic spectrum and its relevance to optical microscopy

## Slide 4: The Light Spectrum
The Electromagnetic Spectrum.
  The “Optical” spectrum regime covers the range of wavelengths from 10-3 m
  (far-infrared) to 10-8 m
  (ultra-violet).
Image Source: http://www.imt.liu.se/edu/courses/TBMT36/artiklar/intro/6.Spectrum.pdf
Image Source:
http://www.cwa-union.org/issues/osh/articles/microwave-radio-frequency-radiation.html

## Slide 5: Some Definitions
Absorption
  When light passes through an object the intensity is reduced depending upon the color absorbed. Thus the selective absorption of white light produces colored light.
Refraction
  Direction change of a ray of light passing from one transparent medium to another with different optical density. A ray from less to more dense medium is bent perpendicular to the surface, with greater deviation for shorter wavelengths
Diffraction
  Light rays bend around edges - new wavefronts are generated at sharp edges - the smaller the aperture the lower the definition
Dispersion
  Separation of light into its constituent wavelengths when entering a transparent medium - the change of refractive index with wavelength, such as the spectrum produced by a prism or a rainbow

## Slide 6: Refraction & Dispersion
Light is “bent” and the resultant colors separate (dispersion).
Red is least refracted, violet most refracted.
dispersion
Short wavelengths are “bent” more than long wavelengths

## Slide 7: Refraction
.
But it is really here!!
He sees the fish here….

## Slide 8
Absorption
Control
No blue/green light
red filter
B & G absorbed

## Slide 9: Light absorption
NOTE: Microsoft powerpoint changes colors according to whatever style you have loaded, so be aware of this “feature” when looking at colors on slides!!
white light
blue light
red light
green light
B & G absorbed
R & G absorbed
B & R absorbed

## Slide 10: Absorption Chart
Color in white light
Color of light absorbed
red
blue
green
magenta
cyan
yellow
blue
blue
blue
blue
green
green
green
green
red
red
red
red
black
gray
green
blue
pink
NOTE: Microsoft powerpoint changes colors according to whatever style you have loaded, so be aware of this “feature” when looking at colors on slides!!

## Slide 11: The light spectrum
Wavelength ---- Frequency
Blue light
488 nm
short wavelength
high frequency
high energy (2 times the red)
Red light
650 nm
long wavelength
low frequency
low energy
Photon as a wave packet of energy

## Slide 12: Magnification
An object can be focussed generally no closer than 250 mm from the eye (depending upon how old you are!)
This is considered to be the normal viewing distance for 1x magnification
Young people may be able to focus as close as 125 mm so they can magnify as much as 2x because the image covers a larger part of the retina - that is it is “magnified” at the place where the image is formed

## Slide 13: Magnification
1000mm
35 mm slide
24x35 mm
M =
1000 mm
35 mm
= 28
        The projected image is 28 times larger than we would see it at 250 mm from our eyes.
If we used a 10x magnifier we would have a magnification of 280x, but we would reduce the field of view by a factor of 10x.
There used to be things called “Slide Projectors”!!!
Look it up on Google!!!
p

## Slide 14: Some Principles of Magnification
Rule of thumb is is not to exceed 1,000 times the NA of the objective
Modern microscopes magnify both in the objective and the ocular and thus are called “compound microscopes” - Simple microscopes have only a single lens

## Slide 15: Basic Microscopy
Bright field illumination does not reveal differences in brightness between structural details - i.e. no contrast
Structural details emerge via phase differences and by staining of components
The edge effects (diffraction, refraction, reflection) produce contrast and detail

## Slide 16: Microscope Basics
© J.Paul Robinson
Originally conformed to the German DIN standard
  Deutsche Industrie Norm, or DIN standard configuration
Standard required the following
  real image formed at a tube length of 160mm
  the parfocal distance set to 45 mm
  object to image distance set to 195 mm
Currently we use the ISO standards
And of course most microscopes are now infinity not 160mm
DIN standard eyepieces have an international standard 23mm diameter
Some objectives have a color ring to help identify the magnification: black (1x), brown (2x), red (4x), yellow (10x), green (20x), turquoise (25x), light blue (40x), dark blue (60x), white (100x).
Focal length
of objective
= 45 mm
Mechanical
tube length
= 160 mm
Object to
Image
Distance
= 195 mm

## Slide 17: The Conventional Microscope
Focal length
of objective
= 45 mm
Object to
Image
Distance
= 195 mm
Mechanical
tube length
= 160 mm
Modified from “Pawley “Handbook of Confocal Microscopy”, Plenum Press

## Slide 18: Microscopes - then to now
Images from Nikon
promotional materials
© J.Paul Robinson
© J.Paul Robinson
Photos: © J. Paul Robinson

## Slide 19: Upright Scope
Brightfield
Source
Epi-
illumination
Source
Image from Nikon
promotional materials

## Slide 20: Inverted Microscope
Brightfield
Source
Epi-
illumination
Source
Image from Nikon
promotional materials

## Slide 21
Image from Nikon
promotional materials
Typical Inverted Microscope
These days we use modern
Digital cameras not 35 mm !!

## Slide 22: Conventional Finite Optics with Telan system
Sample being imaged
Intermediate Image
Telan Optics
Objective
Other optics
Ocular
45 mm
160 mm
195 mm
Modified from “Pawley “Handbook of Confocal Microscopy”, Plenum Press

## Slide 23: Infinity Optics
Sample being imaged
Primary Image Plane
Objective
Other optics
Ocular
Other optics
Tube Lens
Infinite
Image
Distance
The main advantage of infinity corrected lens systems is the relative insensitivity to additional optics within the tube length. Secondly one can focus by moving the objective and not the specimen (stage)
Modified from “Pawley “Handbook of Confocal Microscopy”, Plenum Press

## Slide 24
Images reproduced from:
http://micro.magnet.fsu.edu/

## Slide 25: Objectives
PLAN-APO-40X 1.30 N.A. 160/0.22
Flat field Apochromat Magnification Numerical Tube Coverglass
Factor
Aperture Length Thickness
 - Infinity
corrected

## Slide 26: “Half” of a Zeiss Microscope objective
Photo taken by the author in Zeiss Museum, Jena

## Slide 27: Objectives
Limit for smallest resolvable distance d between 2 points is (Rayleigh criterion):
d =
1.22

Thus, high NUMERICAL APERTURE is critical for high magnification
In a medium of refractive index n the wavelength gets shorter:n
This defines a “resel” or “resolution element”

## Slide 28: Numerical Aperture
Resolving power is directly related to numerical aperture.
The higher the NA the greater the resolution
Resolving power:
  The ability of an objective to resolve two distinct lines very close together
  NA = n sin μ
  (n=the lowest refractive index between the object and first objective element) (hopefully 1)
  μ is 1/2 the angular aperture of the objective

## Slide 29
m
NA=n(sin m)
Light cone
(n=refractive index)

## Slide 30: Numerical Aperture
The wider the angle the lens is capable of receiving light at, the greater its resolving power
The higher the NA, the shorter the working distance
Images reproduced from:
http://micro.magnet.fsu.edu/
Please go to this site and do the tutorials

## Slide 31: Numerical Aperture
For a narrow light beam (i.e. closed illumination aperture diaphragm) the finest resolution is (at the brightest point of the visible spectrum i.e. 530 nm)…(closed condenser).
NA
2 x NA

.00053
2 x 1.00
= 0.265 m
.00053
1.00
= 0.53 m
With a cone of light filling the entire aperture the theoretical resolution is…(fully open condenser)..

=
=

## Slide 32: Raleigh Criterion
Rayleigh criterion is the generally accepted criterion for the minimum resolvable detail - the imaging process is said to be diffraction-limited when the first diffraction minimum of the image of one source point coincides with the maximum of another.
http://hyperphysics.phy-astr.gsu.edu/hbase/phyopt/Raylei.html
By Spencer Bliven - Own work, Public Domain, https://commons.wikimedia.org/w/index.php?curid=31456019
θ
= 1.220
λ
___
D
where θ is the angular resolution (radians), λ is the wavelength of light, and D is the diameter of the lens' aperture.
The factor 1.220 is derived from a calculation of the position of the first dark circular ring surrounding the central Airy disc of the diffraction pattern. This number is more precisely 1.21966989... (OEIS: A245461), the first zero of the order-one Bessel function.
https://en.wikipedia.org/wiki/Angular_resolution

## Slide 33: Object Resolution
Example:
40 x 1.3 N.A. objective at 530 nm light
2 x NA

.00053
2 x 1.3
= 0.20 m
=
40 x 0.65 N.A. objective at 530 nm light
2 x NA

.00053
2 x .65
= 0.405 m
=
R=l/(2NA) 1
R=0.61 l/NA 2
R=1.22 l/(NA(obj) + NA(cond)) 3

## Slide 34
Images reproduced from:
http://micro.magnet.fsu.edu/
Please go to this site and do the tutorials

## Slide 35: Microscope Objectives
Specimen
Coverslip
Oil
Microscope
Objective
Stage
60x 1.4 NA
PlanApo
Standard Coverglass Thickness
#00 = 0.060 - 0.08
#0 = 0.080 - 0.120
#1 = 0.130 - 0.170
#1.5 = 0.160 - 0.190
#2 = 0.170 - 0.250
#3 = 0.280 - 0.320
#4 = 0.380 - 0.420
#5 = 0.500 - 0.60 mm
Cover slips are not all equal!!!

## Slide 36
Refractive Index
Objective
n=1.52
n = 1.52
n = 1.52
Specimen
Coverslip
Oil
n=1.33
n = 1.52
n = 1.0
n = 1.5
Water
n=1.52
Air
Matched RI
Mis-Matched RI

## Slide 37: Summary Lecture 2
Simple versus compound microscopes
Achromatic aberration
Spherical aberration
Refraction, absorption, dispersion, diffraction
Magnification
Upright and inverted microscopes
Optical Designs - 160 mm and infinity optics
Microscope Objectives
Refractive Index